## December 2010
UGC Certified Systems - 2021
Machine | Software Package / Developer | Model | Probe | Frame Grabber | Gain Settings/ Magnification
Aloka 500 - Old Version | BIA/Designer Genes Technologies, Inc. | 1.0 | 5044 -17cm | CX100 | 1.5X magnification, Gain = 90, near gain = -25, far gain = 2.1, Frame Corr = On/Auto, AGC = 1, Contrast = 4, Focus 1 and 2 = On, Focus 3 and 4 = Off
| BIA/Designer Genes Technologies, Inc. | 1 | 5044 -17cm | PXC200 | 1.5X magnification, Gain = 90, near gain = -25, far gain = 2.1, Frame Corr = On/Auto, AGC = 1, Contrast = 4, Focus 1 and 2 = On, Focus 3 and 4 = Off
**Field Technician must also be UGC Lab certified in order to use this system.** | BIA PRO PLUS/Designer Genes Technologies, Inc. | 1 | 5044-17cm | PXC200 | 1.5X magnification, Gain = 90, near gain = -25, Frame Corr = On/Auto, AGC=1, Contrast=4, Focus 1 and 2 = On, Focus 3 and 4 = Off
**Field Technician must also be UGC Lab certified in order to use this system.** | BIA PRO PLUS/Designer Genes Technologies, Inc. | 1 | 5044-17cm | USB | 1.5X magnification, Gain = 90, near gain = -25, Frame Corr = On/Auto, AGC=1, Contrast=4, Focus 1 and 2 = On, Focus 3 and 4 = Off
| Iowa State University | 1.0 | 5044 -17cm | CX100 | 1.5X magnification, Gain = 90, near gain = -25, far gain = 2.1, Frame Corr = On/Auto, AGC = 1, Contrast = 4, Focus 1 and 2 = On, Focus 3 and 4 = Off
| Iowa State University | 1 | 5044 -17cm | PXC200 | 1.5X magnification, Gain = 90, near gain = -25, far gain = 2.1, Frame Corr = On/Auto, AGC = 1, Contrast = 4, Focus 1 and 2 = On, Focus 3 and 4 = Off Filter=off
| UltraInsights / CPEC | 1.0 | 5044 -17cm | CX100 | 1.5X magnification, Gain = 90, near gain = -25, far gain = 2.1, Frame Corr = On/Auto, AGC = 1, Contrast = 4, Focus 1 and 2 = On, Focus 3 and 4 = Off
| UltraInsights / CPEC | 1.0 | 5044 -17cm | LinXcel Video Grabber USB 2.0 | 1.5X magnification, Gain = 90, near gain = -25, far gain = 2.1, Frame Corr = On/Auto, AGC = 1, Contrast = 4, Focus 1 and 2 = On, Focus 3 and 4 = Off
| UltraInsights / CPEC | 1 | 5044 -17cm | PXC200 | 1.5X magnification, Gain = 90, near gain = -25, far gain = 2.1, Frame Corr = On/Auto, AGC = 1, Contrast = 4, Focus 1 and 2 = On, Focus 3 and 4 = Off Filter=off
| The CUP Lab® LLC | 1.0 | 5044 -17cm | CX100 | 1.5X magnification, Gain = 90, near gain = -25, far gain = 2.1, Frame Corr = On/Auto, AGC = 1, Contrast = 4, Focus 1 and 2 = On, Focus 3 and 4 = Off
| The CUP Lab® LLC | 1 | 5044 -17cm | PXC200 | 1.5X magnification, Gain = 90, near gain = -25, far gain = 2.1, Frame Corr = On/Auto, AGC = 1, Contrast = 4, Focus 1 and 2 = On, Focus 3 and 4 = Off
| The CUP Lab® LLC | 1 | 5044 -17cm | USB | 1.5X magnification, Gain = 90, near gain = -25, far gain = 2.1, Frame Corr = On/Auto, AGC = 1, Contrast = 4, Focus 1 and 2 = On, Focus 3 and 4 = Off
| The CUP Lab® LLC | 1 | 5044 -17cm | VCE | 1.5X magnification, Gain = 90, near gain = -25, far gain = 2.1, Frame Corr = On/Auto, AGC = 1, Contrast = 4, Focus 1 and 2 = On, Focus 3 and 4 = Off
Aloka 500 - New Version | BIA/Designer Genes Technologies, Inc. | 1.0 | 5044 -17cm | CX100 | 1.5X magnification, Gain = 90, near gain = -25, far gain = 2.1, Frame Corr = On/Auto, AGC = 1, Contrast = 4, Focus 1 and 2 = On, Focus 3 and 4 = Off
| BIA/Designer Genes Technologies, Inc. | 1 | 5044 -17cm | PXC200 | 1.5X magnification, Gain = 90, near gain = -25, far gain = 2.1, Frame Corr = On/Auto, AGC = 1, Contrast = 4, Focus 1 and 2 = On, Focus 3 and 4 = Off
**Field Technician must also be UGC Lab certified in order to use this system.** | BIA PRO PLUS/Designer Genes Technologies, Inc. | 1 | 5044-17cm | PXC200 | 1.5X magnification, Gain = 90, near gain = -25, Frame Corr = On/Auto, AGC=1, Contrast=4, Focus 1 and 2 = On, Focus 3 and 4 = Off
**Field Technician must also be UGC Lab certified in order to use this system.** | BIA PRO PLUS/Designer Genes Technologies, Inc. | 1 | 5044-17cm | USB | 1.5X magnification, Gain = 90, near gain = -25, Frame Corr = On/Auto, AGC=1, Contrast=4, Focus 1 and 2 = On, Focus 3 and 4 = Off
| Biotronics | 1.0 | 5044 -17cm | CX100 | 2.0X magnification, Gain = 90, near gain = -20, far gain = 2.1, Frame Corr = On/Auto, AGC = 1, Contrast = 4, Focus 1 and 2 = On, Focus 3 and 4 = Off
| Biotronics / BioSoft | 1 | 5044 -17cm | VCE | 2.0X magnification, Gain 90, near gain -20, far gain 2.1, Focus 1, 2=On, Focus 3, 4=Off
| UltraInsights / CPEC | 1.0 | 5044 -17cm | CX100 | 1.5X magnification, Gain = 90, near gain = -25, far gain = 2.1, Frame Corr = On/Auto, AGC = 1, Contrast = 4, Focus 1 and 2 = On, Focus 3 and 4 = Off
| UltraInsights / CPEC | 1.0 | 5044 -17cm | LinXcel Video Grabber USB 2.0 | 1.5X magnification, Gain = 90, near gain = -25, far gain = 2.1, Frame Corr = On/Auto, AGC = 1, Contrast = 4, Focus 1 and 2 = On, Focus 3 and 4 = Off
| UltraInsights / CPEC | 1 | 5044 -17cm | PXC200 | 1.5X magnification, Gain = 90, near gain = -25, far gain = 2.1, Frame Corr = On/Auto, AGC = 1, Contrast = 4, Focus 1 and 2 = On, Focus 3 and 4 = Off Filter=off
| UltraInsights / CPEC | 1 | 5044 -17cm | Avermedia USB | 1.5X magnification, Gain = 90, near gain = -25, far gain = 2.1, Frame Corr = On/Auto, AGC = 1, Contrast = 4, Focus 1 and 2 = On, Focus 3 and 4 = Off Filter=off
| The CUP Lab® LLC | 1.0 | 5044 -17cm | CX100 | 1.5X magnification, Gain = 90, near gain = -25, far gain = 2.1, Frame Corr = On/Auto, AGC = 1, Contrast = 4, Focus 1 and 2 = On, Focus 3 and 4 = Off
| The CUP Lab® LLC | 1 | 5044 -17cm | PXC200 | 1.5X magnification, Gain = 90, near gain = -25, far gain = 2.1, Frame Corr = On/Auto, AGC = 1, Contrast = 4, Focus 1 and 2 = On, Focus 3 and 4 = Off
| The CUP Lab® LLC | 1 | 5044 -17cm | USB | 1.5X magnification, Gain = 90, near gain = -25, far gain = 2.1, Frame Corr = On/Auto, AGC = 1, Contrast = 4, Focus 1 and 2 = On, Focus 3 and 4 = Off
| The CUP Lab® LLC | 1 | 5044 -17cm | VCE | 1.5X magnification, Gain = 90, near gain = -25, far gain = 2.1, Frame Corr = On/Auto, AGC = 1, Contrast = 4, Focus 1 and 2 = On, Focus 3 and 4 = Off
| The CUP Lab® LLC | 1 | 5044 -17cm | Elgato | 1.5X magnification, Gain = 90, near gain = -25, far gain = 2.1, Frame Corr = On/Auto, AGC = 1, Contrast = 4, Focus 1 and 2 = On, Focus 3 and 4 = Off
Pie Medical (Classic) Scanner 200 | Iowa State University | 1.0 | 18cm | CX100 | Machine calibrated with phantom prior to scanning, Pfats = 8 frames per second (FPS), REA & Rump images 5 FPS
| Iowa State University | 1 | 18cm | PXC200 | Machine calibrated with phantom prior to scanning, Pfats = 8 frames per second (FPS), REA & Rump images 5 FPS
Pie Medical (Classic) Scanner 200 | Biotronics / BioSoft | 1 | 5044 -17cm | VCE | Machine calibrated with phantom prior to scanning, IMF images = 8 frames per second
| The CUP Lab® LLC | 1.0 | 18cm | PXC200 | Machine calibrated with phantom prior to scanning, Pfats = 8 frames per second (FPS), REA & Rump images 5 FPS
| The CUP Lab® LLC | 2.0 | 18cm | PXC200 | Machine calibrated with phantom prior to scanning, Pfats = 8 frames per second (FPS), REA & Rump images 5 FPS
| The CUP Lab®LLC | 1 | 18cm | CX100 | Machine calibrated with phantom prior to scanning, Pfats = 8 frames per second (FPS), REA & Rump images 5 FPS
| The CUP Lab® LLC | 1 | 18cm | USB | Machine calibrated with phantom at 8 FPS prior to and after scanning as normal, however all images including Pfats, REA, and Rump images at 5 FPS. None other than phantom calibration images taken at 8 FPS.
Aquila | The CUP Lab® LLC | 1 | 18cm | USB | Total gain= 225 (+/-2), Near Gain = 100 (+/-2), Far Gain = 100 (+/-2), Frame rate is 17 fps.
Aquila | The CUP Lab® LLC | 1 | 18cm | Elgato | Total gain= 225 (+/-2), Near Gain = 100 (+/-2), Far Gain = 100 (+/-2), Frame rate is 17 fps.
SonoVet 2000 | Biotronics | 1.0 | L2-5 170mm | CX100 | IMF images: 12cm (depth setting), N=10, F=1, G=20, FA=1, DR=66, RL=3; Rump and Ribeye cross-sectional images: 16cm (depth setting)
Falco 100 | Biotronics | 1.0 |  | CX100 | IMF images: 26 fps (frames per second), Overall gain at highest level, Near gain at the middle of the range, Far gain at the lowest of the range; Rump and Ribeye cross-sectional images: 17 fps (frames per second)
Evo | UltraInsights | 1 | 17cm | On-board Internal | Pre-set upon machine setup. Technician must select "UltraIn" prior to scanning.
| The CUP Lab® LLC | 1 | 17cm | On-board Internal | Pre-set upon machine setup. Technician must select "CUP LAB" prior to scanning.
Ibex Pro | The CUP Lab® LLC | 1 | 18cm | Elgato | Pre-set upon machine setup.
Exago | The CUP Lab® LLC | 1 | 18cm | LAN connection | Pre-set upon machine setup. 20cm depth
Exago | The CUP Lab® LLC | 1 | 18cm | LAN connection | Pre-set upon machine setup. 16cm depth